IIIT BANGALORE

Water Meter Test Bench

Complete Data Flow — Lab & Bench System

Authoritative reference for all data flow, hardware interaction,
state machine execution, and inter-building communication.

Document Version: 1.0  |  Date: February 2025

Classification: CONFIDENTIAL

ACMIS Technologies LLP

Maxsense Division  |  Hyderabad, India

# 1. System Overview

The IIIT Bangalore Water Meter Test Bench is a two-building system designed for ISO 4064 compliant calibration of DN15/DN20/DN25 water meters using the gravimetric method.

| Aspect | Bench System (Test Bench Building) | Lab System (Lab Building) |
| --- | --- | --- |
| Role | Standalone primary — controls all hardware, runs tests, operates independently | Convenience/monitoring layer — initiates tests remotely, receives results, manages certificates |
| Hardware Control | Full (VFD, valves, sensors, PID, safety) | NONE — never directly controls hardware |
| Connectivity | None required (standalone) | LoRa link to Bench (graceful degradation if down) |
| Django Settings | config.settings_bench | config.settings_lab |
| RPi5 Unit | B1 | L3 |
| Database | db_bench.sqlite3 (primary) | db_lab.sqlite3 (mirror) |
| Real-time | WebSocket via Django Channels (200ms) | HTMX polling (2s) |
| Communication | Via B4 LinkMaster (LoRa 865MHz ASP) | Via L1 LinkMaster (LoRa 865MHz ASP) |

# 2. Physical Architecture

╔═════════════════════════════════════════════════════════════════════╗ ║ LAB BUILDING ║ ║ ║ ║ ┌────────────────────────────────────────────────────┐ ║ ║ │ L3: RPi5 -- Lab Server │ ║ ║ │ ┌──────────────────┐ ┌─────────────┐ │ ║ ║ │ │ Django App │ │ SQLite DB │ │ ║ ║ │ │ (Lab Portal) │ │ db_lab.sqlite│ │ ║ ║ │ │ settings_lab.py │ └─────────────┘ │ ║ ║ │ └────────┬─────────┘ │ ║ ║ │ │ Ethernet LAN --> Lab staff browsers │ ║ ║ │ │ USB Serial │ ║ ║ │ ┌────────┴─────────┐ │ ║ ║ │ │ L2: ESP32 Bridge │ Indoor. USB<->RS485. │ ║ ║ │ └────────┬─────────┘ │ ║ ║ │ │ RS485 │ ║ ║ │ ┌────────┴─────────┐ │ ║ ║ │ │ L1: ESP32+RA01SH │ Rooftop. LoRa<->RS485. │ ║ ║ │ │ Lab LinkMaster │ SX1262, 865MHz IN865 band. │ ║ ║ │ └────────┬─────────┘ │ ║ ║ └───────────┼─────────────────────────────────────────┘ ║ ║ │ ║ ╚══════════════╬══════════════════════════════════════════════════════╝ │ LoRa 865 MHz -- ASP Encrypted (AES-256-CBC + HMAC-SHA256) ~~~~~~~~│~~~~~~~~ Air Gap (200m+) ~~~~~~~~~~~~~~~~~~~~~~~~ │ ╔══════════════╬══════════════════════════════════════════════════════╗ ║ │ TEST BENCH BUILDING ║ ║ ┌───────────┼─────────────────────────────────────────┐ ║ ║ │ ┌────────┴──────────┐ B1: RPi5 -- Bench Controller │ ║ ║ │ │ B4: ESP32+RA01SH │ ┌────────────────────────────┐ │ ║ ║ │ │ Bench LinkMaster │ │ Django App (Bench) │ │ ║ ║ │ │ Rooftop. SX1262. │ │ settings_bench.py │ │ ║ ║ │ └────────┬──────────┘ │ Test Engine + PID + Safety │ │ ║ ║ │ │ RS485->USB │ db_bench.sqlite3 │ │ ║ ║ │ └──────────-> │ 7-inch HDMI Touch LCD (kiosk)│ │ ║ ║ │ └──────┬────────┬────────────┘ │ ║ ║ │ │ │ │ ║ ║ │ USB-1│ │USB-2 │ ║ ║ │ │ │ │ ║ ║ │ ┌──────────────────────────────┴┐ ┌───┴────────────────┐│ ║ ║ │ │ B2: ESP32 -- Sensor Bridge │ │ B3: ESP32 -- VFD ││ ║ ║ │ │ RS485 Bus 1 + GPIOs │ │ Bridge ││ ║ ║ │ │ (EM, Scale, 4-20mA, DUT) │ │ RS485 Bus 2 ││ ║ ║ │ │ (5 valves, tower, E-stop) │ │ (VFD only) ││ ║ ║ │ └──────────┬────────────────────┘ └───┬────────────────┘│ ║ ║ └─────────────┼───────────────────────────┼─────────────────┘ ║ ║ │ │ ║ ║ RS485 Bus 1 RS485 Bus 2 ║ ║ (9600 baud) (9600 baud, isolated) ║ ║ │ │ ║ ║ ┌───────────┼───────────┐ │ ║ ║ │ │ │ │ │ │ ║ ║ F1 F2 F3 F4x2 DUT F5 ║ ║ EM Scale 4-20mA Press Meter VFD Delta ║ ║ addr=1 addr=2 addr=3 addr=20 addr=1 ║ ║ (or manual) ║ ║ ║ ║ B2 GPIOs: Hardwired (NO software): ║ ║ +- BV1: Ball Valve DN25 +- E-Stop Button (NC -> Contactor) ║ ║ +- BV2: Ball Valve DN20 +- Contactor (415V + 24V rail) ║ ║ +- BV3: Ball Valve DN15 +- Exhaust Fan (direct MCB) ║ ║ +- DV1: 3-Way Diverter ║ ║ +- DRN: Drain Valve F6: Pump Kirloskar 3HP ║ ║ +- Tower Light (R/Y/G) (3-phase via VFD, no software) ║ ║ +- E-Stop Monitor (input) ║ ║ ║ ╚════════════════════════════════════════════════════════════════════════╝

# 3. Unit Reference

| ID | Unit | Location | Role |
| --- | --- | --- | --- |
| L1 | Lab LinkMaster | Lab rooftop | ESP32+RA01SH (SX1262, 865MHz). LoRa↔RS485 gateway. |
| L2 | Lab RS485 Bridge | Lab indoor | ESP32. USB↔RS485 transparent bridge to L1. |
| L3 | Lab RPi5 | Lab indoor | Django web portal. Accessed via LAN by lab staff. |
| B1 | Bench RPi5 + Touch LCD | Bench indoor | Main controller. Django + test engine + 7-inch HDMI kiosk. 2× USB to B2, B3. |
| B2 | Sensor Bridge ESP32 | Bench | RS485 Bus 1 (EM, scale, 4-20mA, DUT) + GPIO (5 valves, tower light, E-stop monitor). |
| B3 | VFD Bridge ESP32 | Bench | RS485 Bus 2 (VFD Delta only). Electrically isolated bus. |
| B4 | Bench LinkMaster | Bench rooftop | ESP32+RA01SH (SX1262, 865MHz). LoRa↔RS485 gateway. |

# 4. Field Devices

| ID | Device | Interface | Connected To | Bus / Notes |
| --- | --- | --- | --- | --- |
| F1 | EM Flow Meter DN25 (±0.5%) | RS485 Modbus | B2 | Bus 1, addr=1 |
| F2 | Weighing Scale 200 kg | RS485 | B2 | Bus 1, addr=2 |
| F3 | 4-20mA Modbus Module | RS485 Modbus | B2 | Bus 1, addr=3. Reads 2× pressure + temp. |
| F4 | Pressure Transmitter ×2 | 4-20mA analog | F3 inputs | Upstream (Ch1) + Downstream (Ch2) |
| F5 | VFD Delta VFD022EL43A | RS485 Modbus | B3 | Bus 2, addr=1. Isolated. |
| F6 | Pump Kirloskar 3HP | 3-phase via VFD | F5 | No software interface. |
| BV1 | Ball Valve DN25 | Relay | B2 GPIO | Line select — DN25 meters |
| BV2 | Ball Valve DN20 | Relay | B2 GPIO | Line select — DN20 meters |
| BV3 | Ball Valve DN15 | Relay | B2 GPIO | Line select — DN15 meters |
| DV1 | 3-Way Diverter Valve | Relay pair | B2 GPIO | +/- polarity: collection vs bypass |
| DRN | Drain Valve | Relay | B2 GPIO | Collection tank drain |
| — | Tower Light (R/Y/G) | Digital out | B2 GPIO | 3 outputs for Red/Yellow/Green |
| — | E-Stop Button | Hardwired NC | Contactor coil | NO software in the loop |
| — | Contactor | Hardwired | 415V + 24V rail | E-stop breaks coil → all power cut |
| — | Exhaust Fan | Direct MCB | Contactor/MCB | No software control |
| DUT | Device Under Test | RS485 or Manual | B2 Bus 1 / Touch UI | addr=20 (configurable), or manual entry |

# 5. Scenario 1 — Lab Initiates Test Remotely

| # | Actor / Component | Action |
| --- | --- | --- |
| 1 | Lab Technician | Logs into Lab Portal (http://<L3-ip>:8080) via browser on LAN. |
| 2 | Lab Portal (L3) | Technician registers meter (serial, DN15/20/25, class, type, DUT mode) → saved to Lab SQLite DB. |
| 3 | Lab Portal (L3) | Technician creates new test for the meter. System auto-populates Q1–Q8 from ISO 4064 fixture. Test record created: status='pending', source='lab'. |
| 4 | Lab Django | Packages test request as ASP message: {command: START_TEST, meter_serial, meter_size, meter_class, dut_mode, q_points: [Q1..Q8 params]}. Encrypts with AES-256-CBC, signs with HMAC-SHA256, adds sequence number + timestamp. |
| 5 | L3 → L2 → L1 | ASP frame sent via USB Serial to L2 bridge, forwarded via RS485 to L1 LinkMaster. |
| 6 | L1 → LoRa → B4 | L1 transmits via LoRa 865MHz. If payload > ~200 bytes after encryption, ASP fragments into multiple LoRa packets (max 255 bytes each). |
| 7 | B4 → Bridge → B1 | B4 receives LoRa, forwards via RS485/USB to Bench Django on B1. |
| 8 | Bench Django (B1) | Verifies HMAC signature. Decrypts AES payload. Checks sequence number (replay protection). Creates/mirrors Test record in Bench DB: source='lab', status='pending'. |
| 9 | Bench Django (B1) | Sends ACK back: B1 → B4 → LoRa → L1 → L2 → L3. |
| 10 | Lab Portal (L3) | Receives ACK. Updates test status='acknowledged'. Shows 'Test sent to bench' in UI. |
| 11 | Bench | Test enters the execution queue. If bench is IDLE, begins test execution (Scenario 3). If busy, queues with status='queued'. |

## ASP Message Frame Structure

┌─────────────┬───────────┬────────────┬──────────────────────────┬───────────────┐ │ Device ID │ Seq # │ Timestamp │ AES-256-CBC Encrypted │ HMAC-SHA256 │ │ (4 bytes) │ (2 bytes) │ (4 bytes) │ Payload (variable) │ (32 bytes) │ │ 0x0001=Lab │ Big-endian│ Unix epoch │ IV prepended (16 bytes) │ Over entire │ │ 0x0002=Bench│ │ │ + compressed JSON │ frame │ └─────────────┴───────────┴────────────┴──────────────────────────┴───────────────┘

## Fragmentation (payloads > 200 bytes)

Fragment Header (3 bytes): [Fragment ID (1)] [Fragment Index (1)] [Total Fragments (1)] Each LoRa packet: [Fragment Header (3)] + [Fragment Data (up to 252 bytes)] Receiver reassembles fragments by Fragment ID, validates completeness, then processes. Retry: If ACK not received within 3 seconds, retransmit. Max 3 retries.

# 6. Scenario 2 — Bench Runs Test Locally

| # | Actor / Component | Action |
| --- | --- | --- |
| 1 | Bench Technician | Interacts with 7-inch touch LCD (Chromium kiosk on B1). Auto-logged in as system user (session timeout = 1 year). |
| 2 | Bench Touch UI | Technician registers meter or selects existing one. Saved to Bench SQLite DB. |
| 3 | Bench Touch UI | Technician creates test: selects meter → selects DUT mode (RS485 or manual) → reviews auto-populated Q1–Q8 → confirms. Test record: source='bench', status='pending'. |
| 4 | Bench Touch UI | Technician taps 'START TEST'. Test enters execution engine (Scenario 3). |
| 5 | Bench Django | If LoRa link to lab is active, sends status update: 'Test started locally'. Lab gets notification. If LoRa is down, bench proceeds without lab — no dependency. |

# 7. Scenario 3 — Test Execution: Full State Machine

This is the heart of the system. The Bench Django test engine (controller/state_machine.py) runs as an asyncio task on B1. It orchestrates all hardware through B2 (sensors + GPIO) and B3 (VFD).

## 7.0 State Machine Overview

IDLE → PRE_CHECK → LINE_SELECT → PUMP_START → [Q-point loop: FLOW_STABILIZE → TARE_SCALE → MEASURE → CALCULATE → DRAIN → NEXT_POINT] → COMPLETE EMERGENCY_STOP: reachable from ANY state. ERROR: reachable from any active state (recoverable — operator can retry).

┌──────────────────────────────────────────────────────────────┐ │ │ │ ┌──────┐ ┌───────────┐ ┌─────────────┐ ┌──────────┐ │ │ │ IDLE │───>│ PRE_CHECK │───>│ LINE_SELECT │───>│PUMP_START│ │ │ └──────┘ └───────────┘ └─────────────┘ └────┬─────┘ │ │ ^ │ │ │ │ v │ │ │ ┌─────────────────────────────────────────────┐ │ │ │ │ Q-POINT LOOP (Q1 -> Q8) │ │ │ │ │ │ │ │ │ │ ┌────────────────┐ ┌────────────┐ │ │ │ │ │ │FLOW_STABILIZE │───>│ TARE_SCALE │ │ │ │ │ │ └────────────────┘ └─────┬──────┘ │ │ │ │ │ │ │ │ │ │ │ v │ │ │ │ │ ┌────────────┐ ┌───────────────┐ │ │ │ │ │ │ NEXT_POINT │<──│ DRAIN │ │ │ │ │ │ └──────┬─────┘ └──────^────────┘ │ │ │ │ │ │ │ │ │ │ │ │ next Q? ┌─────┴──────┐ │ │ │ │ │ ¦ yes->loop │ CALCULATE │ │ │ │ │ │ ¦ no v └──────^──────┘ │ │ │ │ │ ¦ │ │ │ │ │ │ ¦ ┌───────┴──────┐ │ │ │ │ │ ¦ │ MEASURE │ │ │ │ │ │ ¦ │ (3 sub-phases)│ │ │ │ │ │ ¦ └───────────────┘ │ │ │ │ └────┼─────────────────────────────────────────┘ │ │ │ v │ │ │ ┌──────────┐ │ │ └───│ COMPLETE │ │ │ └──────────┘ │ │ │ │ ┌─────────────────┐ <- reachable from ANY state │ │ │ EMERGENCY_STOP │ │ │ └─────────────────┘ │ │ ┌─────────────────┐ <- recoverable, operator retry │ │ │ ERROR │ │ │ └─────────────────┘ │ └──────────────────────────────────────────────────────────────┘

## 7.1 State: IDLE

| Aspect | Detail |
| --- | --- |
| Entry | System ready. Pump off, all valves closed, diverter at BYPASS position. |
| Trigger | Test initiated from Touch UI (Scenario 2) or received from Lab via LoRa (Scenario 1). |
| Action | Load ISO 4064 Q1–Q8 parameters for the selected meter size and class. Create TestResult placeholder records for all 8 Q-points. Update test.status='running', test.started_at=now(). |
| Tower Light | GREEN steady |
| Exit | → PRE_CHECK |

## 7.2 State: PRE_CHECK

| Aspect | Detail |
| --- | --- |
| Purpose | Validate all systems before starting. Prevent unsafe test starts. |
| Hardware Reads | B1→B2: Poll Bus 1 (EM meter status, scale status, 4-20mA module status). B1→B3: Poll Bus 2 (VFD status, fault code). B2 GPIOs: Read E-stop monitor (contactor aux contact). |
| Tower Light | YELLOW blink during checks. GREEN flash if all pass. RED if any fail. |
| Exit (pass) | All checks pass → LINE_SELECT |
| Exit (fail) | Show error on touch UI with specific message, return to IDLE |

Safety Checks Performed

| # | Check | Threshold | Fail Action |
| --- | --- | --- | --- |
| 1 | Contactor closed (power available) | Aux contact GPIO = HIGH | BLOCK. 'Power off. Check E-stop and contactor.' |
| 2 | E-stop not active | GPIO = HIGH (NC circuit) | BLOCK. 'E-stop is pressed. Release and reset.' |
| 3 | EM meter responding | Modbus addr 1 responds | BLOCK. 'EM flow meter offline. Check Bus 1.' |
| 4 | Weighing scale responding | Modbus addr 2 responds | BLOCK. 'Weighing scale offline. Check Bus 1.' |
| 5 | 4-20mA module responding | Modbus addr 3 responds | BLOCK. 'Pressure module offline. Check Bus 1.' |
| 6 | VFD responding, no faults | Modbus addr 1, fault=0 | BLOCK. 'VFD fault. Check VFD panel.' |
| 7 | Upstream pressure < max | < 8.0 bar (SAFETY_PRESSURE_MAX) | BLOCK. 'High pressure. Check system.' |
| 8 | Reservoir level > min | > 20% (SAFETY_RESERVOIR_MIN) | BLOCK. 'Low reservoir. Refill before testing.' |
| 9 | Scale weight reasonable | < 180 kg (SAFETY_SCALE_MAX) | BLOCK. 'Scale overloaded. Empty collection tank.' |
| 10 | Temperature in range | 5–40°C (SAFETY_TEMP_MIN/MAX) | BLOCK. 'Temperature out of range.' |

## 7.3 State: LINE_SELECT

| Aspect | Detail |
| --- | --- |
| Purpose | Open the correct ball valve for the DUT meter size. |
| Hardware | B1 → B2 GPIO: Close BV1, BV2, BV3 (all off first). Then open one: DN25→BV1, DN20→BV2, DN15→BV3. |
| Interlock | Mutual exclusion: only one BVx open at a time. Software enforces. Wait for GPIO position feedback. |
| Diverter | DV1 confirmed in BYPASS position (water returns to reservoir during startup). |
| Timeout | 5 seconds (SAFETY_VALVE_TIMEOUT). If no position feedback → ERROR: 'Valve jam. Check mechanical.' |
| Exit | → PUMP_START |

## 7.4 State: PUMP_START

| Aspect | Detail |
| --- | --- |
| Hardware | B1 → B3 → Bus 2 → VFD (addr=1): Write run command (register 0x2000=0x0001 Run Forward). Write initial frequency = 10 Hz (low start, register 0x2001). |
| Confirmation | Read VFD status register. Wait for 'running' flag. Read actual frequency from register 0x2103 ramping up. |
| Timeout | 10 seconds. If VFD doesn't confirm running → ERROR. |
| Tower Light | YELLOW steady (test in progress from here until COMPLETE). |
| Exit | → FLOW_STABILIZE (begins Q-point loop, starting with Q1) |

## 7.5 Q-Point Loop (repeats Q1 through Q8)

For each Q-point, the following 5 states execute in sequence. The pump stays running between Q-points (no restart). Only the VFD frequency changes via PID to hit the new target flow rate.

7.5.1 State: FLOW_STABILIZE

Contains two sub-phases: FLOW_RAMP and FLOW_STABLE.

Sub-phase A — FLOW_RAMP (PID Control Loop): Every 200ms (PID_SAMPLE_RATE):

1. READ: B1 -> B2 -> RS485 Bus 1 -> F1 (EM Meter, addr=1) -> actual flow rate (L/h) 2. COMPARE: actual_flow vs target_flow (from ISO 4064 for current Q-point) 3. PID CALCULATE: error = target_flow - actual_flow P = Kp x error (Kp = 0.5) I = Ki x integral(error x dt) (Ki = 0.1) D = Kd x d(error)/dt (Kd = 0.05) pid_output = P + I + D vfd_freq = clamp(pid_output, 5.0, 50.0) (PID_OUTPUT_MIN, PID_OUTPUT_MAX) 4. WRITE: B1 -> B3 -> RS485 Bus 2 -> F5 (VFD, addr=1) -> set frequency (reg 0x2001) 5. ALSO READ (for monitoring/logging): - F3 (addr=3): upstream pressure, downstream pressure, temperature - F2 (addr=2): current scale weight 6. STORE: SensorReading record in Bench DB (time-series, 5Hz) 7. BROADCAST: - Bench: Redis -> Django Channels -> WebSocket -> Touch LCD gauges (200ms) - Lab: Every 5s, send TEST_STATUS via LoRa ASP: {q_point, flow, pressure, state}

Sub-phase B — FLOW_STABLE: Wait for 5 consecutive readings (1 second) where |actual_flow - target_flow| / target_flow × 100 ≤ 2.0% (SAFETY_FLOW_STABILITY = 2%, SAFETY_STABILITY_COUNT = 5). Once stable, proceed.

Timeout: 60 seconds. If stability not achieved → ERROR with current process value. Operator can adjust PID gains or retry.

Exit: → TARE_SCALE

7.5.2 State: TARE_SCALE

| Aspect | Detail |
| --- | --- |
| Purpose | Zero the weighing scale before collecting water. Critical for gravimetric accuracy. |
| Hardware | B1 → B2 → RS485 Bus 1 → F2 (Scale, addr=2): Send tare command. Read weight until stable at 0.000 ±0.020 kg. |
| Record | Store tare_weight and tare_timestamp. |
| Timeout | 5 seconds. If scale doesn't zero → ERROR: 'Scale tare failed.' |
| Exit | → MEASURE |

7.5.3 State: MEASURE

Contains three sub-phases: DIVERT_OPEN, COLLECTING, and DIVERT_CLOSE.

Sub-phase A — DIVERT_OPEN:

1. Record start values: - DUT totalizer: RS485 read from addr=20 (if DUT mode='rs485') OR prompt manual BEFORE reading on touch LCD (if DUT mode='manual') Manual entry: large numeric keypad, validation (reasonable range) - EM meter totalizer: RS485 read from F1 (addr=1) 2. Switch diverter: B1 -> B2 GPIO -> DV1 relay pair -> polarity for COLLECTION position (Water now flows into collection tank on weighing scale) 3. Record: divert_start_timestamp

Sub-phase B — COLLECTING:

Water flows through DUT into collection tank sitting on weighing scale. Every 200ms (PID continues running): - F2 (Scale, addr=2): read current weight - F1 (EM Meter, addr=1): read flow rate + totalizer - F3 (4-20mA, addr=3): read pressures + temperature - PID continues maintaining target flow rate Target weight calculation: target_weight_kg = target_volume_L x water_density(temperature_C) (Density lookup table per ISO 4064, e.g., 20C -> 0.99820 kg/L) Continue until: current_weight >= target_weight -> proceed to DIVERT_CLOSE

Sub-phase C — DIVERT_CLOSE:

1. Switch DV1 back to BYPASS position (Water now returns to reservoir via bypass) 2. Wait 2 seconds for flow to settle in collection tank 3. Record final values (stabilized): - F2 (Scale): final_weight (wait for scale reading to stabilize +/-0.010 kg) - F1 (EM Meter): final_totalizer - DUT: final_totalizer (RS485 read from addr=20) OR prompt manual AFTER reading on touch LCD (if DUT mode='manual') Manual validation: after_reading > before_reading, reasonable range 4. Record: divert_end_timestamp

Exit: → CALCULATE

7.5.4 State: CALCULATE

| Calculation | Formula |
| --- | --- |
| Reference Volume | ref_volume_L = (final_weight - tare_weight) / water_density(temperature_C). Density from ISO 4064 lookup table. Example: 20°C → 0.99820 kg/L, 25°C → 0.99705 kg/L |
| DUT Volume | dut_volume_L = final_dut_totalizer - start_dut_totalizer. For manual mode: after_reading - before_reading |
| Error % | error_pct = ((dut_volume - ref_volume) / ref_volume) × 100 |
| Pass/Fail | abs(error_pct) ≤ mpe_pct → PASS, else → FAIL. MPE from ISO 4064: ±5% for Q1 (lower zone), ±2% for Q2–Q4+ (upper zone) |

Store: Save TestResult record with all values:

TestResult( test=test, q_point='Q3', # Current Q-point target_flow_lph=100.0, # From ISO 4064 fixture actual_flow_lph=99.5, # Average flow during collection ref_volume_l=10.050, # Gravimetric reference (weight/density) dut_volume_l=10.120, # DUT totalizer difference error_pct=0.696, # ((10.120-10.050)/10.050)x100 mpe_pct=2.0, # From ISO 4064 (upper zone) passed=True, # ¦0.696¦ <= 2.0 pressure_up_bar=2.5, # Upstream pressure during test pressure_dn_bar=2.3, # Downstream pressure during test temperature_c=22.1, # Water temperature during test duration_s=360, # Seconds from divert_open to divert_close weight_kg=10.032, # Net weight collected zone='Upper', # ISO 4064 zone classification )

Send to Lab: If LoRa active, send individual Q-point TEST_RESULT via ASP. Lab stores and updates live monitor.

Exit: → DRAIN

7.5.5 State: DRAIN

| Aspect | Detail |
| --- | --- |
| Hardware | B1 → B2 GPIO: Open DRN (drain valve). Monitor scale weight via F2. Wait for weight to return to near-tare (±0.050 kg). Close DRN. |
| Duration | Depends on collected volume. Small Q-points (Q1: ~1L) drain fast. Large Q-points (Q4: ~100L) take longer. |
| Timeout | 120 seconds. If weight doesn't return → ERROR: 'Drain timeout. Check drain valve.' |
| Exit | → NEXT_POINT |

7.5.6 State: NEXT_POINT

| Aspect | Detail |
| --- | --- |
| Logic | If current Q-point < Q8: increment to next Q-point. Pump stays running (no restart). → FLOW_STABILIZE with new target flow rate. If current Q-point = Q8: all points complete. → COMPLETE. |
| UI Update | Touch LCD: Q-point stepper advances. Green checkmark on completed points (pass) or red × (fail). Lab: updated via LoRa. |

## 7.6 State: COMPLETE

| Aspect | Detail |
| --- | --- |
| Pump Stop | B1 → B3 → VFD: Write stop command (register 0x2000=0x0005). Wait for VFD stopped status. |
| Valves | Close all ball valves (BV1, BV2, BV3). DV1 to BYPASS. DRN closed. |
| Calculate Overall | overall_pass = all 8 TestResult.passed are True. Update Test record: status='completed', overall_pass=T/F, completed_at=now(). |
| Tower Light | PASS: GREEN blink 3× then steady. FAIL: RED blink 3× then steady. |
| Touch UI | Show verdict banner: green 'PASSED' or red 'FAILED'. Navigate to Results tab with Q-point result cards. |
| Lab Notification | If LoRa active: send TEST_COMPLETE message with overall_pass and summary of all 8 Q-points. |
| Certificate | If passed and operator has permission: generate certificate PDF. Otherwise pending manager approval (lab side). |
| Exit | → IDLE. System ready for next test. |

## 7.7 State: EMERGENCY_STOP

| Aspect | Detail |
| --- | --- |
| Triggers | (a) Hardware E-stop: contactor drops (B2 GPIO detects aux contact change). (b) Software: touch UI ABORT button, safety watchdog interlock violation, lab EMERGENCY_STOP command via LoRa. |
| Hardware E-Stop | Hardwired. Contactor coil breaks. VFD + pump + 24V rail lose power instantly. Valves spring-return to closed. NO SOFTWARE INVOLVED. B2 GPIO just detects it happened after the fact. |
| Software E-Stop | B1→B3: VFD emergency stop command. B1→B2 GPIO: close all valves, DV1 to bypass. Set test.status='aborted'. Log event with timestamp and reason. |
| Safety Watchdog | Runs every 200ms. Triggers software E-stop if: pressure > 8.0 bar, scale > 180 kg, temp outside 5–40°C, reservoir < 20%, any ESP32 communication timeout (>2 seconds). |
| Tower Light | RED steady. |
| Touch UI | Red blinking banner: 'EMERGENCY STOP ACTIVE'. Requires manual reset from touch UI to return to IDLE. |
| Lab Notify | If LoRa active: send EMERGENCY_STOP notification with reason code. |

## 7.8 State: ERROR (Recoverable)

| Aspect | Detail |
| --- | --- |
| Triggers | Non-critical failures: valve jam, sensor timeout, VFD fault (non-dangerous), scale tare failure, flow stability timeout, drain timeout. |
| Action | Pump remains at current state (not stopped unless dangerous). Touch UI shows orange error banner with specific message and 'Retry' / 'Abort' buttons. |
| Retry | Operator taps 'Retry' → re-enters the failed state. Up to 3 retries per state. |
| Abort | Operator taps 'Abort' → EMERGENCY_STOP sequence → IDLE. |
| Tower Light | YELLOW blink. |

# 8. Scenario 4 — Lab Receives Test Results

| # | Component | Action |
| --- | --- | --- |
| 1 | Bench Django | After each Q-point CALCULATE: packages individual TestResult as ASP message. Sends via B4 → LoRa → L1 → L2 → L3. |
| 2 | Lab Django (L3) | Receives each Q-point result. Decrypts, verifies. Stores TestResult in Lab DB. Updates dashboard (HTMX polling 2s picks up changes). |
| 3 | Bench Django | After Q8 COMPLETE: sends TEST_COMPLETE summary (overall_pass, all 8 points summary, timestamps). |
| 4 | Lab Django | Receives TEST_COMPLETE. Updates Test record: status='completed'. Sets approval_status='pending'. |
| 5 | Lab Portal | Dashboard shows notification badge. Manager sees 'Pending Approval' in queue. |
| 6 | Lab Manager | Reviews Q1–Q8 results in Lab Portal. Views error curve chart (error % vs Q-point). Clicks Approve or Reject (with comments). |
| 7 | Lab Django | If approved: generates certificate PDF (reports app). Assigns certificate number (format: IIITB-YYYYMMDD-NNNN). Available for download. Sends approval status back to bench via LoRa. |
| 8 | Bench Django | Receives approval. Updates local Test record. Certificate available on bench too. |

Selective Retransmission

If the lab misses any Q-point result (detected by gap in sequence), it sends a RESULT_REQUEST message specifying the missing Q-point numbers. Bench retransmits only those specific results. 3 retry attempts before marking COMM_FAILURE on that Q-point (operator can manually sync later).

# 9. Scenario 5 — Emergency Stop

## 9.1 Hardware E-Stop (Primary — No Software)

Physical mushroom-head E-stop button ¦ +-- NC (Normally Closed) contact in series with contactor coil ¦ +-- Contactor drops out ¦ +-- 415V cut -> VFD loses power -> Pump stops immediately +-- 24V cut -> Valve solenoids de-energize -> Spring-return to CLOSED +-- Exhaust fan stops ¦ +-- 5V rail STAYS powered (RPi5, ESP32s, Touch LCD remain on) ¦ +-- B2 GPIO detects contactor aux contact -> LOW ¦ +-- Software logs event, updates UI 'E-STOP ACTIVE', sets test.status='aborted', notifies Lab via LoRa. But the shutdown already happened electrically.

## 9.2 Software E-Stop (Secondary)

Touch UI 'ABORT' button ¦ +-- WebSocket -> Django Channels -> Test Engine ¦ +-- Test Engine executes: +-- B1->B3->Bus 2->VFD: Emergency stop command (reg 0x2000=0x0003) +-- B1->B2 GPIO: Close BV1, BV2, BV3, DRN +-- B1->B2 GPIO: DV1 to BYPASS +-- Set test.status='aborted', log reason='Operator abort' +-- Notify Lab via LoRa (if active)

## 9.3 Safety Watchdog Auto-Trigger

Runs every 200ms during any active test state. Triggers software E-stop if:

| Condition | Threshold | Reason Code |
| --- | --- | --- |
| Upstream pressure too high | > 8.0 bar | PRESSURE_HIGH |
| Scale overloaded | > 180 kg | SCALE_OVERLOAD |
| Temperature too low | < 5°C | TEMP_LOW |
| Temperature too high | > 40°C | TEMP_HIGH |
| Reservoir level too low | < 20% | RESERVOIR_LOW |
| B2 communication timeout | > 2 seconds no response | B2_COMM_TIMEOUT |
| B3 communication timeout | > 2 seconds no response | B3_COMM_TIMEOUT |
| Contactor aux contact lost | GPIO goes LOW | POWER_LOST |

## 9.4 Lab-Initiated Emergency Stop

Lab Portal -> 'EMERGENCY STOP' button (available to ALL roles) ¦ +-- Lab Django -> ASP message {command: EMERGENCY_STOP, reason: 'Lab operator'} -> L3->L2->L1 -> LoRa -> B4->B1 ¦ +-- Bench Django treats same as local software E-stop: VFD stop, valves close, test aborted.

# 10. Scenario 6 — Real-Time Monitoring

## 10.1 Bench Side (WebSocket — True Real-Time)

During an active test, every PID cycle (200ms), the sensor manager publishes readings to a Redis channel. Django Channels WebSocket consumer pushes to the touch LCD browser.

Test Engine (asyncio task) ¦ every 200ms +-- Read all sensors (B2 Bus 1 + B3 Bus 2) +-- Store SensorReading in DB (5Hz time-series) +-- Publish to Redis channel: 'test.{test_id}.sensors' ¦ +-- Django Channels Consumer ¦ +-- WebSocket push to Touch LCD browser ¦ +-- Alpine.js updates gauges with smooth needle animation: +-- Flow rate gauge (current value + target line) +-- Pressure gauge (upstream bar + downstream bar) +-- Temperature display (C) +-- VFD frequency display (Hz) +-- PID output display (Hz) +-- Scale weight (real-time accumulation, kg) +-- Q-point stepper (Q1 check Q2 check Q3 active Q4 pending ... Q8) +-- State indicator (FLOW_STABILIZE / COLLECTING / etc.)

## 10.2 Lab Side (HTMX Polling — Near Real-Time)

Lab does NOT get 200ms data. LoRa bandwidth limits updates to periodic summaries.

Lab Portal Browser ¦ +-- HTMX polls every 2 seconds: ¦ GET /api/tests/{id}/status/ ¦ +-- Lab Django responds with latest data received via LoRa: ¦ { ¦ "status": "running", ¦ "current_q_point": "Q3", ¦ "current_state": "COLLECTING", ¦ "flow_rate_lph": 100.2, ¦ "pressure_up_bar": 2.5, ¦ "temperature_c": 22.1, ¦ "weight_kg": 5.230, ¦ "updated_at": "2026-02-13T14:30:22Z" ¦ } ¦ +-- Browser updates: +-- Test status badge (Running / Completed / etc.) +-- Current Q-point indicator +-- Current state name +-- Basic sensor readings (updated every ~5s via LoRa)

Bench sends TEST_STATUS via LoRa every 5 seconds during active test.

## 10.3 Tower Light States

| System State | Tower Light | Touch LCD Status Strip |
| --- | --- | --- |
| IDLE / Ready | GREEN steady | Green: 'System Ready' |
| PRE_CHECK running | YELLOW blink | Yellow: 'Running pre-checks...' |
| PRE_CHECK failed | RED steady (3s) → off | Red: 'Pre-check failed: [reason]' |
| Test Running (Q-loop) | YELLOW steady | Blue: 'TEST RUNNING — Q3 Flow Stabilize' |
| COMPLETE — PASS | GREEN blink 3× → steady | Green: 'TEST COMPLETE — PASSED' |
| COMPLETE — FAIL | RED blink 3× → steady | Red: 'TEST COMPLETE — FAILED' |
| ERROR (recoverable) | YELLOW blink | Orange: 'ERROR — [message]. Tap to retry.' |
| EMERGENCY_STOP | RED steady | Red blink: 'EMERGENCY STOP ACTIVE' |

# 11. Communication Protocol Reference

## 11.1 USB Serial (RPi5 ↔ ESP32)

| Parameter | Value |
| --- | --- |
| Baud rate | 115200 |
| Format | 8N1 |
| Protocol | JSON command/response |

Command format

{"cmd": "MB_READ", "bus": 1, "addr": 1, "reg": 0, "count": 2} {"cmd": "MB_WRITE", "bus": 1, "addr": 1, "reg": 0, "value": 100} {"cmd": "GPIO_SET", "pin": "BV1", "state": 1} {"cmd": "GPIO_GET", "pin": "BV1"} {"cmd": "VALVE", "valve": "BV1", "action": "OPEN"} {"cmd": "DIVERTER", "position": "COLLECT"} {"cmd": "STATUS"}

Response format

{"ok": true, "data": {"value": 123.45}} {"ok": false, "error": "TIMEOUT", "message": "No response from addr 1"}

## 11.2 RS485 Bus 1 (B2 ↔ Sensors + DUT)

| Parameter | Value |
| --- | --- |
| Baud rate | 9600 |
| Protocol | Modbus RTU |
| Parity | None (8N1) |

| Device | Address | Key Registers |
| --- | --- | --- |
| EM Flow Meter | 1 | Flow rate (L/h), Totalizer (L), Status |
| Weighing Scale | 2 | Weight (kg), Tare command, Status |
| 4-20mA Module | 3 | Ch1: Upstream pressure, Ch2: Downstream pressure, Ch3: Temperature |
| DUT (meter under test) | 20 (configurable) | Totalizer (L) |

## 11.3 RS485 Bus 2 (B3 ↔ VFD)

| Parameter | Value |
| --- | --- |
| Baud rate | 9600 |
| Protocol | Modbus RTU |
| Isolation | Electrically isolated bus (VFD noise protection) |

| Register | R/W | Description |
| --- | --- | --- |
| 0x2000: Control word | Write | 0x0001=Run, 0x0003=E-Stop, 0x0005=Stop |
| 0x2001: Frequency setpoint | Write | Hz × 100 (e.g. 3000 = 30.00 Hz) |
| 0x2103: Actual output frequency | Read | Hz × 100 |
| 0x2104: Actual output current | Read | A × 100 |
| 0x2100: Status word | Read | Running, Fault, At Speed flags |
| 0x2105: Fault code | Read | Last fault code |

## 11.4 LoRa (L1 ↔ B4 via ASP)

| Parameter | Value |
| --- | --- |
| Frequency | 865 MHz (IN865 band for India) |
| Chip | SX1262 (RA-01SH module) |
| Spreading Factor | SF10 (configurable SF7–SF12) |
| Bandwidth | 125 kHz |
| Coding Rate | 4/5 |
| TX Power | +22 dBm |
| Preamble | 8 symbols |
| Max LoRa payload | 255 bytes |
| Encryption | AES-256-CBC (16-byte IV prepended) |
| Authentication | HMAC-SHA256 (32-byte tag appended) |
| Replay protection | Sequence number (2 bytes, monotonically increasing) |
| Fragmentation | Payloads > ~200 bytes split into fragments |
| ACK timeout | 3 seconds |
| Max retries | 3 |

## 11.5 B2 GPIO Pin Map

| Function | Type | Notes |
| --- | --- | --- |
| BV1 (DN25 valve) | Digital OUT (relay) | Line select |
| BV2 (DN20 valve) | Digital OUT (relay) | Line select |
| BV3 (DN15 valve) | Digital OUT (relay) | Line select |
| DV1+ (diverter +) | Digital OUT (relay) | Polarity pair for 3-way |
| DV1- (diverter -) | Digital OUT (relay) | Polarity pair for 3-way |
| DRN (drain valve) | Digital OUT (relay) | Collection tank drain |
| Tower RED | Digital OUT | Tower light |
| Tower YELLOW | Digital OUT | Tower light |
| Tower GREEN | Digital OUT | Tower light |
| E-Stop Monitor | Digital IN | Contactor aux contact (NC) |
| BV1 Position | Digital IN (optional) | Valve feedback |
| BV2 Position | Digital IN (optional) | Valve feedback |
| BV3 Position | Digital IN (optional) | Valve feedback |

GPIO Pins: GPIO pin numbers TBD — will be assigned during B2 firmware development.

# 12. Data Storage & Sync

| Data | Bench DB | Lab DB | Sync Method |
| --- | --- | --- | --- |
| TestMeter registration | Yes (primary) | Yes (mirror) | LoRa ASP (meter info included in START_TEST request) |
| Test record | Yes (primary) | Yes (mirror) | LoRa ASP (status updates + completion) |
| TestResult (Q-points) | Yes (primary) | Yes (mirror) | LoRa ASP (per Q-point result + completion summary) |
| ISO4064Standard | Pre-loaded fixture | Pre-loaded fixture | Same JSON fixture loaded on both sides |
| SensorReading (5Hz) | Yes (time-series) | No | Not synced. Bench-only. Used for PID tuning + diagnostics. |
| Certificate PDF | Generated after approval | Generated after approval | Both can generate independently once approval is synced |
| User accounts | Independent | Independent | Not synced. Each side manages own users. |
| Audit log | Independent | Independent | Each side logs own actions. |

# 13. LoRa Message Types

| Message Type | Direction | Trigger | Payload Summary |
| --- | --- | --- | --- |
| START_TEST | Lab → Bench | Technician creates test | meter info, Q-point params, DUT mode |
| START_TEST_ACK | Bench → Lab | Bench receives START_TEST | test_id, status=acknowledged |
| TEST_STATUS | Bench → Lab | Every 5s during test | q_point, state, flow, pressure, temp |
| TEST_RESULT | Bench → Lab | After each Q-point | q_point, error_pct, passed, all measurements |
| TEST_COMPLETE | Bench → Lab | After Q8 done | overall_pass, summary of all 8 points |
| RESULT_REQUEST | Lab → Bench | Missing Q-point detected | list of missing Q-point numbers |
| EMERGENCY_STOP | Lab → Bench | Lab operator presses abort | reason string |
| EMERGENCY_ACK | Bench → Lab | Bench processes E-stop | status=aborted, reason |
| APPROVAL_STATUS | Lab → Bench | Manager approves/rejects | approval_status, certificate_number |
| HEARTBEAT | Bidirectional | Every 30s when idle | device_id, uptime, status |

# 14. Key Design Principles

1. Bench is standalone: Operates with zero lab connectivity. All hardware control is local to bench RPi5. Lab is a convenience layer, not a dependency.

2. Lab never controls hardware: Lab only sends requests (START_TEST, EMERGENCY_STOP) and receives results. No direct valve/pump/sensor commands ever traverse the LoRa link.

3. Gravimetric reference: Primary reference volume measured by weight (200 kg scale), not by EM flow meter. EM meter is for PID feedback and cross-check only.

4. ISO 4064 compliance: All 8 Q-points (Q1–Q8), MPE values (±5% lower zone, ±2% upper zone), water density correction by temperature, test procedures follow the standard.

5. Defense in depth: Hardware E-stop (electrical, no software in loop) → Software safety watchdog (200ms) → PID output clamping (5–50 Hz) → Valve mutual exclusion interlocks.

6. Encrypted communication: All LoRa traffic AES-256-CBC encrypted, HMAC-SHA256 authenticated, sequence-numbered for replay protection.

7. Single codebase, dual deployment: One Django project with shared apps (accounts, meters, testing, comms, reports, audit). Side-specific apps: controller + bench_ui (bench only), lab_ui (lab only). Configured via settings_bench.py / settings_lab.py.

8. Graceful degradation: If LoRa link goes down, bench continues testing without interruption. Results are queued and synced when link recovers.

END OF DOCUMENT

ACMIS Technologies LLP | Maxsense Division | Confidential