# LAB SIDE — Hardware Group Mapping

## L1: LoRa LinkMaster (Rooftop)

Purpose

Receives ASP protocol frames from L2 via RS485, transmits over LoRa 433MHz to bench building B4. Receives LoRa packets from B4, forwards to L2 via RS485. Pure gateway — no data processing or storage.

Components

| Part | Model/Spec | Connection & Role |
| --- | --- | --- |
| ESP32 | WROOM-32 | Main MCU. Runs LoRa+RS485 bridge firmware. |
| LoRa Module | RA01SH (SX1262) | SPI connection to ESP32. 433MHz, configurable SF/BW/CR. |
| RS485 Transceiver | MAX485 | UART2 on ESP32. Half-duplex. Connected to L2 via 2-wire+GND cable. |
| Antenna | 433MHz omni, SMA | External antenna on mast for line-of-sight. |
| Power | 12V DC input | Voltage regulator to 3.3V for ESP32. Weather-protected supply. |
| Enclosure | IP65 ABS/metal | Rooftop mounted. Cable glands for RS485 + power. |

Firmware Functions

1) Initialize LoRa (SX1262 via SPI, set frequency/SF/BW). 2) Initialize RS485 (UART2, 9600 baud). 3) Main loop: if RS485 data available, read frame, transmit via LoRa. If LoRa packet received, forward to RS485. 4) No encryption/decryption — ASP frames pass through opaquely. 5) LED blink on TX/RX for status.

## L2: RS485-USB Bridge (Indoor)

Purpose

Connects RPi5 (USB) to L1 (RS485). Transparent bridge — ASP protocol frames from Django serial handler pass through to RS485 and vice versa. No data processing.

Components

| Part | Model/Spec | Connection & Role |
| --- | --- | --- |
| ESP32 | WROOM-32 | Main MCU. Bridge firmware. |
| USB-TTL | CP2102 or CH340 | USB to RPi5. Appears as /dev/ttyUSBx. 115200 baud. |
| RS485 Transceiver | MAX485 | UART2. Half-duplex. 9600 baud to L1 rooftop. |

Firmware Functions

1) Initialize USB-UART (115200 baud). 2) Initialize RS485-UART (9600 baud). 3) Main loop: if USB data, forward to RS485. If RS485 data, forward to USB. 4) Frame detection: look for STX/ETX bytes to maintain frame integrity. 5) No encryption, no Modbus — pure byte relay.

## L3: Raspberry Pi 5 (Lab Server)

Purpose

Runs the Django web portal accessible on local LAN. Lab staff create tests, monitor progress, view results, generate certificates — all via browser. Communicates with bench via L2→L1→LoRa→B4→bench.

Software Components

| Component | Detail |
| --- | --- |
| Django Web Portal | accounts, meters, testing, comms, dashboard, reports, audit apps |
| Serial Handler | Background thread. pyserial on /dev/ttyUSB0 (to L2). ASP frame encode/decode. |
| Protocol Library | comms/protocol.py + crypto.py. AES-256, HMAC, sequence numbers. |
| Database | SQLite (dev) or SQLite (prod — single user sufficient for lab). |
| Web Server | Django runserver (dev) or Gunicorn+Nginx (prod). Port 8080. |
| Network | Ethernet on LAN. Users access via http://<ip>:8080 from their laptops/tablets. |

# TEST BENCH SIDE — Hardware Group Mapping

## B1: Raspberry Pi 5 + 7-inch Touch LCD (Bench Controller)

Purpose

The brain of the test bench. Runs the complete test engine: state machine, PID controller, sensor polling, valve control, gravimetric calculation, safety watchdog, certificate generation, and the touch LCD kiosk UI. This is the PRIMARY system — it works fully without lab connection.

Hardware

| Component | Detail |
| --- | --- |
| RPi5 | 4/8GB RAM. 32GB+ microSD (A2 class). |
| Touch LCD | 7-inch capacitive, HDMI connection. 1024x600. Already tested and working. |
| USB Port 1 | Connected to B2 (Sensor Bus Bridge). /dev/ttyBENCH_BUS via udev rule. |
| USB Port 2 | Connected to B3 (VFD Bus Bridge). /dev/ttyVFD_BUS via udev rule. |
| USB Port 3 | Available. Could connect LoRa bridge directly or reserve for debug. |
| Power | 5V/5A USB-C PSU. On separate 5V rail (stays powered during E-stop). |
| OS | Raspberry Pi OS Lite 64-bit. No desktop. Openbox + Chromium kiosk. |

Software Components

| Module | Function |
| --- | --- |
| Django Web App | Same shared project as lab. Bench-specific settings. Kiosk UI templates optimized for 7-inch touch. |
| controller/state_machine.py | Asyncio task. 12-state test execution engine. Coordinates all hardware. |
| controller/pid_controller.py | PID loop: reads EM meter flow (via B2), adjusts VFD freq (via B3). 200ms cycle. |
| controller/sensor_manager.py | Polls B2 for sensor data (EM, scale, pressure, DUT). Polls B3 for VFD status. Stores in Redis. |
| controller/valve_controller.py | Sends GPIO commands to B2. Valve sequencing, mutual exclusion, timing, position feedback. |
| controller/gravimetric.py | Tare, measure, density correction, volume calculation. |
| controller/safety_monitor.py | Parallel watchdog: checks interlocks every 200ms. Can force EMERGENCY_STOP. |
| controller/dut_interface.py | RS485 reader (via B2 bus) + manual entry handler (touch UI prompt). |
| comms/lora_handler.py | ASP protocol handler for lab communication. Runs independently. Bench works without it. |
| PostgreSQL 16 | Production database. Concurrent access for test engine + UI + comms. |
| Redis 7 | Live sensor cache, pub/sub for real-time UI updates, task queue. |

## B2: Sensor Bus Bridge ESP32

Purpose

The most complex ESP32 in the system. Serves THREE functions: (1) RS485 Modbus master for Bus 1 sensors (EM meter, scale, 4-20mA module, DUT), (2) GPIO controller for valves, tower light, and E-stop monitoring, (3) USB bridge to RPi5. RPi5 sends commands, B2 executes them on the bus and GPIOs.

Components

| Part | Model/Spec | Role |
| --- | --- | --- |
| ESP32 | WROOM-32 | MCU. Handles USB, RS485, and GPIO simultaneously. |
| USB-TTL | CP2102/CH340 | USB to RPi5 (115200 baud). Command/response protocol. |
| MAX485 | RS485 transceiver | Bus 1: EM meter, weighing scale, 4-20mA module, DUT. 9600 baud. |
| Relay Module 1 | 4-ch relay | Valves BV1, BV2, BV3, DRN (4 ball valves). |
| Relay Module 2 | 2-ch relay | 3-way diverter (relay A + relay B, polarity based). |
| GPIO Output | Digital pin | Tower light control (R/Y/G — may need 3 GPIOs or single relay-switched). |
| GPIO Input | Digital pin | E-stop / contactor auxiliary contact monitor. Pull-up, NC contact. |

RS485 Bus 1 Devices

| Device | Modbus Addr | Baud | Registers Read/Written |
| --- | --- | --- | --- |
| EM Flow Meter | 1 | 9600 | Flow rate (L/h), totalizer (L), status |
| Weighing Scale 200kg | 2 | 9600 | Weight (kg), tare command, status |
| 4-20mA Input Module | 3 | 9600 | Ch1: upstream pressure (bar), Ch2: downstream pressure (bar) |
| DUT (test meter) | 20 (config) | 9600 (config) | Totalizer (L), flow rate (L/h). Only for RS485-capable meters. |

GPIO Pin Map

| GPIO | Function | Direction | Notes |
| --- | --- | --- | --- |
| TBD | Valve BV1 (DN25) | OUTPUT | Relay: HIGH=open, LOW=close |
| TBD | Valve BV2 (DN20) | OUTPUT | Relay: HIGH=open, LOW=close |
| TBD | Valve BV3 (DN15) | OUTPUT | Relay: HIGH=open, LOW=close |
| TBD | Valve DRN (drain) | OUTPUT | Relay: HIGH=open, LOW=close |
| TBD | Diverter Relay A | OUTPUT | DV1 direction A (+) |
| TBD | Diverter Relay B | OUTPUT | DV1 direction B (+). NEVER both HIGH. |
| TBD | Tower Light Red | OUTPUT | Error / E-stop active |
| TBD | Tower Light Yellow | OUTPUT | Test in progress |
| TBD | Tower Light Green | OUTPUT | System ready / test passed |
| TBD | E-Stop Monitor | INPUT | Contactor aux contact. LOW=E-stop active. Pull-up. |

Pin Assignment: All GPIO pins marked TBD. Will be confirmed from Maxsense board library or custom pin config.

Firmware Functions

B2 runs a command-response protocol over USB with the RPi5. The RPi5 sends commands like: READ_MODBUS(addr, register, count), WRITE_MODBUS(addr, register, value), SET_GPIO(pin, state), GET_GPIO(pin), SET_VALVE(id, state), GET_VALVE_STATUS(). B2 executes and returns results. B2 also has a safety interrupt: if E-stop input goes LOW, immediately de-energize all valve relays and report to RPi5.

## B3: VFD Bus Bridge ESP32

Purpose

Dedicated RS485 bridge for VFD control on isolated Bus 2. Simpler than B2 — only handles Modbus communication with the Delta VFD. Isolated bus prevents VFD electrical noise from affecting sensor readings on Bus 1.

Components

| Part | Model/Spec | Role |
| --- | --- | --- |
| ESP32 | WROOM-32 | MCU. VFD Modbus bridge firmware. |
| USB-TTL | CP2102/CH340 | USB to RPi5 (115200 baud). /dev/ttyVFD_BUS. |
| MAX485 | RS485 transceiver | Bus 2: VFD Delta only. 9600 baud. |

VFD Modbus Registers (Delta VFD022EL43A)

| Register | Name | R/W | Description |
| --- | --- | --- | --- |
| 0x2001 | Frequency Cmd | Write | Target freq ×100 (e.g. 3000 = 30.00 Hz) |
| 0x2000 | Control Word | Write | Bit 0: Run/Stop. Bit 1: Forward/Reverse. |
| 0x2103 | Actual Frequency | Read | Current output freq ×100 |
| 0x2104 | Output Current | Read | Current ×10 (e.g. 150 = 15.0 A) |
| 0x2100 | Status Word | Read | Bit 0: Running. Bit 2: Fault. Bit 3: At speed. |
| 0x2105 | Fault Code | Read | Last fault code (overload, overcurrent, etc.) |

Register Addresses: Delta Modbus registers vary by VFD model. Confirm from VFD022EL43A manual. Above are typical Delta MS300 series registers.

Firmware Functions

Same command-response protocol as B2 over USB. RPi5 sends: VFD_SET_FREQ(hz), VFD_RUN(), VFD_STOP(), VFD_READ_STATUS(). B3 translates to Delta Modbus register writes/reads and returns results. If USB disconnects, B3 sends VFD stop command as safety measure.

## B4: LoRa LinkMaster (Bench Rooftop)

Purpose

Identical to L1 (Lab LinkMaster). Receives LoRa packets from lab L1, forwards to bench via RS485. Sends bench ASP frames over LoRa to lab. Pure gateway.

Components & Firmware

Same as L1: ESP32 + RA01SH (SX1262) + MAX485 + antenna + IP65 + 12V. Firmware is identical to L1 — same binary, different DEVICE_ID configured at flash time. The RS485 side connects to an indoor bridge (could be B2 with a second UART, or a dedicated small bridge ESP32). For simplicity, B4 connects to RPi5 via a third USB-serial adapter or shares B2's RS485 line with address-based routing.

Implementation Decision: B4 LoRa link to RPi5: either (a) third USB-serial ESP32 bridge, or (b) share B2 RS485 bus with protocol-level separation. Option (a) is cleaner. Can use a simple ESP32+MAX485+CP2102 like L2.

## Field Devices Summary

| ID | Device | Qty | Interface | Connected To |
| --- | --- | --- | --- | --- |
| F1 | EM Flow Meter DN25 (±0.5%) | 1 | RS485 Modbus | B2 Bus 1 (addr 1) |
| F2 | Weighing Scale 200kg | 1 | RS485 | B2 Bus 1 (addr 2) |
| F3 | 4-20mA Modbus Module | 1 | RS485 Modbus | B2 Bus 1 (addr 3). Reads 2x pressure. |
| F4 | Pressure Transmitter (0-10bar) | 2 | 4-20mA | F3 module Ch1 (upstream), Ch2 (downstream) |
| F5 | Electric Ball Valves | 5 | Relay | B2 GPIOs (BV1,BV2,BV3,DRN,DV1) |
| F6 | Tower Light (R/Y/G) | 1 | Digital | B2 GPIOs (3 outputs or relay-switched) |
| F7 | VFD Delta 2.2kW | 1 | RS485 Modbus | B3 Bus 2 (dedicated) |
| F8 | Pump Kirloskar 3HP | 1 | 3-phase via VFD | F7 VFD output |
| F9 | Main Contactor | 1 | Hardwired | E-stop breaks coil. Powers VFD+24V. |
| F10 | E-Stop Button | 1 | NC hardwired | Series with F9 contactor coil |
| F11 | Exhaust Fan | 1 | Direct from MCB | No software control |
| F12 | DUT (test meter) | 1 | RS485 or manual | B2 Bus 1 (addr 20) or touch panel entry |

END OF DOCUMENT