# 1. Project Overview for Claude Code

This document provides everything Claude Code needs to build the IIIT-B Water Meter Test Bench software. It is the single reference for project structure, development sequence, and implementation details.

## 1.1 What We're Building

Single Django project that runs on both Lab RPi5 and Bench RPi5 with different settings files. Shared apps for common functionality, side-specific apps for hardware control (bench) and remote monitoring (lab).

## 1.2 Key Constraints

Deadline: March 1, 2025. Three weeks from start.

Hardware not ready: Build with simulators first. All hardware interaction through abstraction layers that can be swapped for real drivers later.

Single developer: Harsha is building everything. Code must be clean, well-commented, easy to debug.

Target hardware: Raspberry Pi 5 (both sides). 7-inch touch LCD on bench. No cloud.

# 2. Django Project Structure

iiitb-testbench/

├── manage.py

├── requirements.txt

├── config/

│   ├── settings_base.py        # Shared settings

│   ├── settings_lab.py         # Lab: DJANGO_SETTINGS_MODULE=config.settings_lab

│   ├── settings_bench.py       # Bench: DJANGO_SETTINGS_MODULE=config.settings_bench

│   ├── urls.py                 # Root URL config

│   ├── wsgi.py

│   └── asgi.py                 # For Django Channels (bench WebSocket)

│

├── accounts/                       # SHARED - User management

│   ├── models.py               # CustomUser with role field

│   ├── views.py                # Login, logout, user CRUD

│   ├── permissions.py          # Role-based decorators/mixins

│   └── templates/accounts/     # Login page, user management

│

├── meters/                         # SHARED - Meter registry

│   ├── models.py               # TestMeter model

│   ├── views.py                # CRUD views

│   └── templates/meters/       # List, add, detail

│

├── testing/                        # SHARED - Test management

│   ├── models.py               # Test, TestResult, ISO4064Standard

│   ├── views.py                # Test creation, results, history

│   ├── services.py             # Business logic (error calc, pass/fail)

│   ├── iso4064.py              # ISO parameter auto-generation

│   └── fixtures/iso4064.json   # DN15/20/25 R160 test parameters

│

├── comms/                          # SHARED - Communication

│   ├── protocol.py             # ASP frame encoder/decoder

│   ├── crypto.py               # AES-256 + HMAC-SHA256

│   ├── serial_handler.py       # pyserial USB handler

│   ├── message_queue.py        # Outgoing queue, ACK tracking, retry

│   └── lora_handler.py         # LoRa-specific message handling

│

├── reports/                        # SHARED - Certificate generation

│   ├── generator.py            # ReportLab PDF builder

│   ├── error_curve.py          # Matplotlib error curve chart

│   └── views.py                # Download endpoints

│

├── audit/                          # SHARED - Audit logging

│   ├── models.py               # AuditLog, CommunicationLog

│   └── middleware.py           # Auto-log all actions

│

├── controller/                     # BENCH ONLY - Hardware control

│   ├── state_machine.py        # 12-state test engine (asyncio)

│   ├── pid_controller.py       # PID flow control

│   ├── sensor_manager.py       # Modbus polling via B2 and B3

│   ├── valve_controller.py     # Valve sequencing via B2 GPIOs

│   ├── gravimetric.py          # Weighing + density correction

│   ├── dut_interface.py        # DUT RS485 + manual entry

│   ├── safety_monitor.py       # Interlock watchdog

│   ├── vfd_controller.py       # VFD commands via B3

│   ├── tower_light.py          # Tower light state mapping

│   └── simulator.py            # Mock hardware for dev/testing

│

├── lab_ui/                         # LAB ONLY - Web portal templates

│   ├── views.py                # Lab-specific views

│   └── templates/lab_ui/       # Sidebar layout, desktop-optimized

│

├── bench_ui/                       # BENCH ONLY - Touch LCD templates

│   ├── views.py                # Touch-optimized views

│   ├── consumers.py            # WebSocket for real-time gauges

│   └── templates/bench_ui/     # Bottom-tab layout, 1024x600

│

├── static/

│   ├── css/                    # Custom SCSS, design tokens

│   ├── js/                     # Alpine.js components, gauge scripts

│   └── img/                    # ACMIS/IIIT-B logos

│

├── templates/

│   ├── base_lab.html           # Lab base: sidebar + header

│   └── base_bench.html         # Bench base: bottom tabs + status strip

│

└── firmware/                       # ESP32 Arduino projects

├── bridge_sensor/              # B2: Sensor bus bridge + GPIO

├── bridge_vfd/                 # B3: VFD bus bridge

├── linkmaster/                 # L1/B4: LoRa LinkMaster

└── bridge_lora/                # L2: Lab LoRa bridge

# 3. Database Models (Django ORM)

## 3.1 accounts/models.py

class CustomUser(AbstractUser):

ROLE_CHOICES = [

('admin', 'Admin'),

('manager', 'Manager'),

('bench_tech', 'Bench Technician'),

('lab_tech', 'Lab Technician'),

]

role = models.CharField(max_length=20, choices=ROLE_CHOICES)

full_name = models.CharField(max_length=200)

# Django's default: username, password, email, is_active, last_login

## 3.2 meters/models.py

class TestMeter(models.Model):

SIZE_CHOICES = [('DN15','DN15'), ('DN20','DN20'), ('DN25','DN25')]

CLASS_CHOICES = [('R80','R80'), ('R100','R100'), ('R160','R160'), ('R200','R200'), ('R315','R315')]

TYPE_CHOICES = [('mechanical','Mechanical'), ('ultrasonic','Ultrasonic'), ('electromagnetic','EM')]

DUT_MODE_CHOICES = [('rs485','RS485 Modbus'), ('manual','Manual Entry')]

serial_number = models.CharField(max_length=50, unique=True)

meter_size = models.CharField(max_length=10, choices=SIZE_CHOICES)

meter_class = models.CharField(max_length=10, choices=CLASS_CHOICES, default='R160')

manufacturer = models.CharField(max_length=100, blank=True)

model = models.CharField(max_length=100, blank=True)

meter_type = models.CharField(max_length=30, choices=TYPE_CHOICES)

dut_mode = models.CharField(max_length=10, choices=DUT_MODE_CHOICES, default='manual')

modbus_address = models.IntegerField(default=20)  # For RS485 mode

modbus_baud = models.IntegerField(default=9600)

register_totalizer = models.IntegerField(default=0)  # Modbus register for volume

registered_by = models.ForeignKey(settings.AUTH_USER_MODEL, on_delete=models.SET_NULL, null=True)

created_at = models.DateTimeField(auto_now_add=True)

## 3.3 testing/models.py

class ISO4064Standard(models.Model):

meter_size = models.CharField(max_length=10)

meter_class = models.CharField(max_length=10)

q_point = models.CharField(max_length=5)  # Q1-Q8

flow_rate_lph = models.FloatField()

test_volume_l = models.FloatField()

duration_s = models.IntegerField()

mpe_pct = models.FloatField()  # Max Permissible Error %

zone = models.CharField(max_length=10)  # Lower/Upper

class Test(models.Model):

STATUS = [('pending','Pending'), ('running','Running'), ('completed','Completed'),

('failed','Failed'), ('aborted','Aborted')]

meter = models.ForeignKey(TestMeter, on_delete=models.CASCADE)

status = models.CharField(max_length=20, choices=STATUS, default='pending')

initiated_by = models.ForeignKey(settings.AUTH_USER_MODEL, on_delete=models.SET_NULL, null=True, related_name='initiated_tests')

approved_by = models.ForeignKey(settings.AUTH_USER_MODEL, on_delete=models.SET_NULL, null=True, blank=True, related_name='approved_tests')

overall_pass = models.BooleanField(null=True)

started_at = models.DateTimeField(null=True)

completed_at = models.DateTimeField(null=True)

created_at = models.DateTimeField(auto_now_add=True)

certificate_pdf = models.CharField(max_length=255, blank=True)

notes = models.TextField(blank=True)

source = models.CharField(max_length=10, default='bench', choices=[('bench','Bench'),('lab','Lab')])

class TestResult(models.Model):

test = models.ForeignKey(Test, on_delete=models.CASCADE, related_name='results')

q_point = models.CharField(max_length=5)

target_flow_lph = models.FloatField()

actual_flow_lph = models.FloatField(null=True)

ref_volume_l = models.FloatField(null=True)       # Gravimetric reference

dut_volume_l = models.FloatField(null=True)        # DUT meter reading

error_pct = models.FloatField(null=True)

mpe_pct = models.FloatField()

passed = models.BooleanField(null=True)

pressure_up_bar = models.FloatField(null=True)

pressure_dn_bar = models.FloatField(null=True)

temperature_c = models.FloatField(null=True)

duration_s = models.IntegerField(null=True)

## 3.4 Bench-Only Models

# controller/models.py (bench only)

class SensorReading(models.Model):

timestamp = models.DateTimeField(auto_now_add=True)

flow_rate_lph = models.FloatField(null=True)

pressure_up = models.FloatField(null=True)

pressure_dn = models.FloatField(null=True)

weight_kg = models.FloatField(null=True)

temperature_c = models.FloatField(null=True)

vfd_freq_hz = models.FloatField(null=True)

vfd_current_a = models.FloatField(null=True)

reservoir_pct = models.FloatField(null=True)

test = models.ForeignKey('testing.Test', null=True, on_delete=models.SET_NULL)

q_point = models.CharField(max_length=5, blank=True)

class DUTManualEntry(models.Model):

test = models.ForeignKey('testing.Test', on_delete=models.CASCADE)

q_point = models.CharField(max_length=5)

reading_before = models.FloatField()

reading_after = models.FloatField(null=True)

entered_by = models.ForeignKey(settings.AUTH_USER_MODEL, on_delete=models.SET_NULL, null=True)

entered_at = models.DateTimeField(auto_now_add=True)

# 4. Development Sequence (Build Order)

Build in this exact order. Each step depends on the previous. Test each step before moving on.

## Phase 1: Foundation (Days 1-5)

| # | Task | Files | Test |
| --- | --- | --- | --- |
| 1 | Create Django project with dual settings | config/ | python manage.py check --settings=config.settings_lab |
| 2 | CustomUser model + migrations | accounts/models.py | Create user, verify role field |
| 3 | Login/logout views + templates | accounts/views.py | Login with each role, verify redirect |
| 4 | Role-based permission decorators | accounts/permissions.py | @role_required('admin') blocks lab_tech |
| 5 | TestMeter model + CRUD views | meters/ | Register meter, list, edit, delete |
| 6 | ISO4064Standard model + fixture | testing/ | python manage.py loaddata iso4064 |
| 7 | Test + TestResult models | testing/models.py | Create test, auto-populate Q1-Q8 |
| 8 | Test creation view with auto Q1-Q8 | testing/views.py | Select meter → see 8 auto Q-points |
| 9 | Base templates (lab sidebar + bench tabs) | templates/ | Visual layout correct on both |

## Phase 2: Communication + Controller (Days 6-12)

| # | Task | Files | Test |
| --- | --- | --- | --- |
| 10 | ASP protocol encoder/decoder | comms/protocol.py | Encode→decode roundtrip, verify HMAC |
| 11 | AES-256 + HMAC crypto module | comms/crypto.py | Encrypt→decrypt, tamper detection |
| 12 | Serial handler (pyserial) | comms/serial_handler.py | Send/receive on /dev/ttyUSB0 |
| 13 | Message queue + ACK tracking | comms/message_queue.py | Send msg, wait ACK, retry on timeout |
| 14 | Hardware simulator (mock all sensors) | controller/simulator.py | Fake flow, weight, pressure, VFD |
| 15 | Sensor manager (Modbus polling) | controller/sensor_manager.py | Poll simulator, get readings |
| 16 | PID controller module | controller/pid_controller.py | Input flow, output freq, verify convergence |
| 17 | Safety watchdog | controller/safety_monitor.py | Trigger on low reservoir, overpressure |
| 18 | Valve controller + sequencing | controller/valve_controller.py | Open BV1, verify BV2/BV3 closed |
| 19 | VFD controller | controller/vfd_controller.py | Set freq, read status via simulator |

## Phase 3: Test Engine + UI + Integration (Days 13-21)

| # | Task | Files | Test |
| --- | --- | --- | --- |
| 20 | State machine (12 states) | controller/state_machine.py | Run full Q1-Q8 cycle on simulator |
| 21 | Gravimetric engine | controller/gravimetric.py | Tare, collect, density correct, calc error |
| 22 | DUT interface (RS485 + manual) | controller/dut_interface.py | Read mock DUT, test manual prompt |
| 23 | Lab web portal full UI (all pages) | lab_ui/ | All 11 pages functional with HTMX |
| 24 | Bench touch UI (all pages) | bench_ui/ | All 7 pages functional for 1024x600 |
| 25 | Live monitor with gauges | bench_ui/ + lab_ui/ | Real-time gauge updates during test |
| 26 | Manual DUT entry overlay | bench_ui/ | Touch keypad, validate before/after |
| 27 | Error curve chart | reports/error_curve.py | Generate chart with Q1-Q8 data |
| 28 | PDF certificate generator | reports/generator.py | Generate cert, verify content |
| 29 | Manager approval workflow | testing/views.py | Approve→generate cert flow |
| 30 | LoRa comm handler | comms/lora_handler.py | Send/receive ASP via LoRa simulator |
| 31 | ESP32 firmware (4 units) | firmware/ | Flash and test each board |
| 32 | Kiosk deployment (bench) | scripts/ | Boot to app, no desktop visible |
| 33 | E2E integration test | -- | Full test cycle: create→run→results→cert |

# 5. Hardware Simulator Specification

Since hardware is not ready, ALL hardware interaction goes through an abstraction layer. The simulator implements the same interface as real hardware drivers. Swap by changing one setting in config.

# config/settings_base.py

HARDWARE_BACKEND = 'simulator'  # or 'real'

# controller/sensor_manager.py

if settings.HARDWARE_BACKEND == 'simulator':

from controller.simulator import SimulatedSensors as SensorBackend

else:

from controller.drivers.real_sensors import RealSensors as SensorBackend

## 5.1 Simulator Behavior

| Sensor | Simulated Behavior |
| --- | --- |
| EM Flow Meter | Returns flow_rate matching VFD frequency with ±3% noise. Totalizer increments based on flow_rate × dt. Responds to Modbus reads. |
| Weighing Scale | Starts at 0 after tare. Increases at rate = flow_rate × density / 3600 when diverter is in collection mode. Returns current weight. Supports tare command. |
| Pressure Upstream | Returns 1.0 + (vfd_freq/50) × 3.0 bar with ±0.1 noise. Higher pressure at higher flow. |
| Pressure Downstream | Returns upstream - 0.2 - (flow_rate/1000) × 0.5 bar (pressure drop across DUT). |
| Temperature | Returns 24.0 ± 0.5 °C (slow random walk). |
| VFD | Accepts frequency command. Ramps to target at 2Hz/sec. Returns actual freq, current = freq×0.3A. Status: running/stopped. |
| Reservoir Level | Starts at 85%. Decreases slowly during test. Resets on drain return. |
| Valves | Instant open/close with 500ms simulated delay. Returns position feedback. |
| DUT (RS485) | Returns totalizer that increments proportionally to flow, with configurable error (e.g. -2% for a slightly under-reading meter). |
| E-Stop / Contactor | Always OK unless explicitly triggered by test scenario. |

# 6. REST API Endpoints

Django views (class-based or function). Both lab and bench expose the same API. Lab uses it for HTMX partial updates. Bench uses it for touch UI + WebSocket.

| Method | URL | Auth | Description |
| --- | --- | --- | --- |
| POST | /api/auth/login/ | Any | Login, return session |
| POST | /api/auth/logout/ | Any | Logout |
| GET | /api/auth/me/ | Any | Current user + role |
| GET | /api/users/ | Admin | List all users |
| POST | /api/users/ | Admin | Create user |
| PATCH | /api/users/{id}/ | Admin | Update user |
| GET | /api/meters/ | Auth | List meters |
| POST | /api/meters/ | Tech+ | Register meter |
| GET | /api/meters/{id}/ | Auth | Meter detail + test history |
| GET | /api/standards/{size}/{class}/ | Auth | ISO 4064 Q1-Q8 params |
| GET | /api/tests/ | Auth | List tests (role-filtered) |
| POST | /api/tests/ | Tech+ | Create test (auto Q1-Q8) |
| GET | /api/tests/{id}/ | Auth | Test detail + results |
| POST | /api/tests/{id}/start/ | Tech+ | Start test (bench: start engine, lab: send to bench) |
| POST | /api/tests/{id}/abort/ | Auth | Abort running test |
| POST | /api/tests/{id}/approve/ | Manager+ | Approve results |
| GET | /api/tests/{id}/results/ | Auth | Q1-Q8 results array |
| GET | /api/reports/{id}/pdf/ | Manager+ | Download certificate PDF |
| GET | /api/system/status/ | Auth | Sensor status, comm health, reservoir level |
| POST | /api/system/emergency-stop/ | Auth | Software E-stop (all roles) |
| GET | /api/sensors/live/ | Auth | Current sensor readings (from Redis cache) |
| GET | /api/audit/ | Admin/Mgr | Audit log entries |
| GET | /api/comms/log/ | Admin | Communication log |

# 7. ESP32 Firmware Specifications

## 7.1 Firmware Variants

| Firmware | Units | Core Function |
| --- | --- | --- |
| bridge_sensor | B2 | USB↔RS485 bridge + GPIO control. Most complex. Handles Modbus forwarding AND direct I/O. |
| bridge_vfd | B3, L2 | Simple USB↔RS485 bridge. Transparent Modbus pass-through. Same firmware for B3 and L2 (just bridge). |
| linkmaster | L1, B4 | RS485↔LoRa bridge. Same firmware, different DEVICE_ID. Uses RA01SH (SX1262) tested library. |

## 7.2 B2 Sensor Bridge — Command Protocol (USB)

RPi5 sends text commands over USB serial (115200 baud). B2 parses and executes. Responses are JSON-like.

// Commands from RPi5 to B2:

MB_READ <addr> <reg> <count>     // Read Modbus registers

MB_WRITE <addr> <reg> <value>    // Write Modbus register

GPIO_SET <pin> <0|1>             // Set GPIO output

GPIO_GET <pin>                   // Read GPIO input

VALVE <id> <OPEN|CLOSE>          // Valve command with safety checks

DIVERTER <COLLECT|BYPASS|HOLD>   // 3-way valve command

STATUS                           // Get all GPIO states + bus health

// Responses from B2:

OK <data>                        // Success with optional data

ERR <code> <message>             // Error with code and description

EVENT ESTOP <0|1>                // Unsolicited: E-stop state change

## 7.3 Key Libraries

| Library | Usage |
| --- | --- |
| RA01SH Library (Harsha's tested) | SX1262 LoRa communication. Pin config in library. Used by L1 and B4. |
| ModbusMaster (Arduino) | Modbus RTU master on B2 and B3 for sensor/VFD communication. |
| SoftwareSerial / HardwareSerial | UART management. USB on Serial0, RS485 on Serial2. |
| ArduinoJson | JSON parsing for command responses (optional, can use simple text). |

# 8. Python Dependencies

# requirements.txt

django==5.0

pyserial==3.5

pycryptodome==3.20

reportlab==4.1

matplotlib==3.8

Pillow==10.2

channels==4.0

daphne==4.1

redis==5.0

psycopg2-binary==2.9  # For PostgreSQL (bench prod)

minimalmodbus==2.1     # Python Modbus RTU (bench, for direct serial if needed)

weasyprint==61.0       # Alternative PDF gen

whitenoise==6.6        # Static file serving

# 9. Key Configuration Values

# config/settings_bench.py (additions to base)

HARDWARE_BACKEND = 'simulator'  # Change to 'real' when hardware ready

BENCH_SERIAL_PORT_BUS1 = '/dev/ttyBENCH_BUS'  # B2 sensor bridge

BENCH_SERIAL_PORT_BUS2 = '/dev/ttyVFD_BUS'     # B3 VFD bridge

BENCH_SERIAL_BAUD = 115200  # USB to ESP32

MODBUS_BAUD = 9600          # RS485 bus

# PID defaults (tuned on-site)

PID_KP = 0.5

PID_KI = 0.1

PID_KD = 0.05

PID_OUTPUT_MIN = 5.0   # Hz

PID_OUTPUT_MAX = 50.0  # Hz

PID_SAMPLE_RATE = 0.2  # seconds (200ms)

# Safety limits

SAFETY_PRESSURE_MAX = 8.0      # bar

SAFETY_RESERVOIR_MIN = 20.0    # percent

SAFETY_SCALE_MAX = 180.0       # kg (200kg scale, 90% limit)

SAFETY_TEMP_MIN = 5.0          # Celsius

SAFETY_TEMP_MAX = 40.0         # Celsius

SAFETY_VALVE_TIMEOUT = 5.0     # seconds

SAFETY_FLOW_STABILITY = 2.0    # percent tolerance

SAFETY_STABILITY_COUNT = 5     # consecutive readings

# ACMIS Protocol

ASP_AES_KEY = 'generate_with_secrets.token_hex(32)'

ASP_HMAC_KEY = 'generate_with_secrets.token_hex(32)'

ASP_DEVICE_ID = 0x0002  # Bench = 0x0002, Lab = 0x0001

END OF DOCUMENT